## BK 879B LCR Meter 092915
Perform BK 879B Setup before starting measurements
Step | Description
1 | Press Power button for 1 sec.
2 | Press L/C/R/Z until 'L' is displayed in upper left corner of display.
3 | Press FREQ button 3 times or until 100 Hz is displayed in upper right corner of display.
4 | Attach test leads to +/- inputs to LCR meter.
5 | Touch test leads together, meter should read < 2.5 mH.
6 | Meter is ready for measurement.
7 | Repeat setup if meter times out between use.
8 | To turn off meter, press Power button for 1 sec when done taking measurements.
BK 879B Calibration Test Procedure
Step | Description
1 | Perform steps 1-3 of Setup procedure.
2 | With nothing connected to the input terminals or test leads separated, press and hold REL/CAL button for 2 sec. LCD should read OPEN in upper right of display.
3 | To calibrate OPEN condition, press REL/CAL button once; system will calibrate and return to the normal display.
4 | With a solid jumper wire between the input terminals or test leads touching, press and hold REL/CAL button for 2 sec. LCD should read SHRT in upper right of display.
5 | To calibrate SHORT condition, press REL/CAL button once; system will calibrate and return to the normal display.
6 | System is ready for normal use.

## Inductance Measurement Std Work
Step | Inductance Measurement Procedure (092916)
1 | Use calibrated BK 879B LCR meter. To setup meter for measurements, use BK Inductance Meter Setup tab in Excel file.
2 | With meter set to measure Inductance, touch test leads together, display should read < 2.5 mH.
3 | Place one test lead on Main solenoid coil spade and the other on the valve body, when reading is steady, record reading. Acceptable range is 350 to 700 mH, nominal is 480 mH.
4 | Place one test lead on Pilot solenoid coil spade and the other on the valve body, when reading is steady, record reading. Acceptable range is 250 to 650 mH, nominal is 380 mH.
BK 879B LCR Meter
Typical measurement setup

## Gas Valve Std Work 051116
Standard Work - Gas Valve Incoming Sample Inspection Process - DRAFT COPY
A Valve shall be rejected and sent to NCM if it does not meet the requirements of this Standard Work
If Valve is Dropped during the Sample Inspection Process, SCRAP VALVE, DO NOT COUNT AS FAIL, Document in Comment Section in Data Sheet.
Step | Procedure | Results
1 | Access the Sample Inspection database, and the "Create PartReceived" screen will come up for you to fill in.
2 | Select Vendor name from the "VendorId" pull-down menu. Tab down
3 | The "SampleInspectionEntryDate" field is automatically filled in. Tab down
4 | Select Auditor Name from the "AuditorId" pull-down menu. Tab down
5 | Scan or type in the Valve "Part Number" information. Tab down
6 | Select "LCINSP" from "Where Found" pull-down menu. Tab down
7 | Select "Sample Inspection" from "Inspection Type" pull-down menu. Tab down
8 | Fill in the "Incoming Date" information (MM/DD/YYYY, e.g., 04/15/2016). (This is the date in which we received the parts from Dexen) Tab down
9 | Determine how many valves need to be tested by using the 3x Inspection Sheet (see 3X Inspection Criteria tab). This sheet is also on the clip board at the work station.
10 | Pull your first box of valves and open it. Pull a valve out and find the 5-digit GM date code (ex: 16035), and record that number into the "Date Code" field. Tab down
11 | There are two boxes for an inspector number. The first box is a required field. The inspector numbers are 3-digit numbers stamped on the valve (ex: 186). Type one number into the first box. If there is a second number tab down and enter that number into that second box. (If no Inspector Number is stamped on the valve, the valve should be rejected and sent to NCM). Tab down
12 | Scan or type the Valve Serial Number. Log valve SN. Tab down
13 | The "Was Tested" box is a default to "Yes". If any of the above information is missing on the valve, the valve failed, and you do not have to test it. If you do not test the valve, change this field to "No". The next two fields are left blank, unless there is a failure. Click on "Create".
14 | Visually inspect Valve for missing / broken screws, debris in inlet / outlet ports/any damage Log results. Tab down | P/F
16 | Measure "Pilot Valve Solenoid Inductance" by placing the red and black leads from the Multimeter on the valve (one of the leads will connect to the valve body and one of the leads will connect to the orange male connector on the valve). Flag if outside 250-650 mH. Log results. Tab down | nom 380 mH
17 | Measure "Main Valve Solenoid Inductance" by placing the red and black leads from the Multimeter on the valve (one of the leads will connect to the valve body and one of the leads will connect to the green male connector on the valve). Fflag if outside 350-700 mH. Log results. Tab down. | nom 480 mH | / ECM / Remote Control (if required)
18 | Install Valve on test bench with gas supply in the "On" position. Use an aim-flame to verify that the valve will not fail open; check at the pilot and main burner. Log results. Tab down | P/F
19 | Verify that the inlet pressure is set to the nominal for each gas type (NG: 7 IWC, LP 12 IWC) If pressures are not meeting the nominal; adjust at the regulator until it does meet nominal.
20 | Turn on Pilot / Burner with Wall Switch or Remote Control.
21 | Verify Pilot Main lights within 60 seconds and burner lights within 10 seconds. Log results. | P/F
22 | Using the aim-flame, test for leaks around the whole valve; front, back and sides. Log results. Tab down | P/F
23 | For Stepper Motor Valves: Adjust flame height on Remote Control to low to get the output pressure Manometer reading. Log results. Tab down. Adjust flame with Remote Control to high and record output pressure Manometer reading. Log results. Tab down. Confirm flame height adjustment visually responds properly for flame height setting. Log results. Tab down | (303/302 specs - High Spec NG 3.5 +/- 0.3 IWC, LP 10.0 +/- 0.5 IWC; Low Spec NG: 1.0 +0.5/-0.1, LP: 2.7 +.9/-.1)
24 | For Manual Valves: Adjust flame height to low by turning the Hi/Lo knob on the Valve to low to get the output pressure Manometer reading. Log results. Tab down. Adjust flame to high by turning the hi/lo knob to high to get the output pressure Manometer reading. Log results. Tab down. Confirm flame height adjustment visually responds properly for flame height setting. Log results. Tab down | 750-500, 501; 593-500,501 - High Spec NG 3.5 +/- 0.3 IWC, LP 10.0 +/-0.5 IWC, Low Spec NG 1.7 +/-0.3, LP 5.5 +/- 0.5
25 | Turn off Pilot / Burner with Wall Switch or Remote Control. Log results. Tab down | P/F
26 | Remove Valve from test bench. If the Valve passes, place black X mark just below the white identification label right above the circle cutout and replace in box. If the Valve fails, click on "Edit This Part Received". This will take you back to the create screen. Click on "Individual Parts Comment" and write in a description of the failure. Tab down and fill in the NCM red tag number. Complete the NCM red tag and send the part to quality for confirmation testing.
27 | When you have completed the testing of all valves, tape the box shut. Take a purple label and place a check mark on it, then place the label on the box. If a valve failed take a colored dot label (not purple) and write the correct qty on it. Place one quantity label on each end of the box.
28 | If the pallet of valves passes, release to production. All non-tested valves still get a purple label but don't put a check mark on the label.
29 | Follow the NCM process for failed valves (reference the Failure Modes tab for lot / pallet rejection conditions).
Example Test Bench Set-Up
Valve In Test Bench
Valve Solenoids, Stepper Motor, Valve In / Out Pressure Measurement Ports
Manometer Connections
3-Digit Dexen Inspector Number
5-Digit Grand Mate Date Code
Valve Serial Number
Valve Part Number

## 3X Inspection Criteria
Incoming Inspection Procedure (3X Inspection)
Function: | Procedure to release new or modified purchased parts to production through incoming inspection
Purpose: | To ensure that parts that meet the Inspection Criteria are inspected prior to being released to production
Inspection Criteria
Parts are placed on Incoming Inspection if the following criteria occurs:
| 1 | If multiple occurrences of NCM material is found
| 2 | A SCAR has been issued to a vender for NCM materials
| 3 | If a revision change is occurring on a part due to an ECO or SPCN
| 4 | If product in transit needs to be Quarantined
| 5 | If a new purchased part is released for a new design on a PRO
Procedure:
Quality Function (Prior to Inspection)
1 | Contact Commodity / Purchasing Representative to Flag specific part number for inspection
2 | Update the Dock Sheet for the Receiving Department of parts that need to be transferred to LC-INSP
3 | Determine the necessary test for the new component (full functionality, visual, or dimensional test)
4 | If the change required reliability approval contact the reliability department for test protocol and equipment needs
Purchasing Function (Prior to Inspection)
1 | Flag the part number for inspection in the Manage 2000 Purchase function - under PO a Y is added to Field 14
If the part is on a new purchase order or under a blanket purchase order
Reliability Function (Prior to Inspection)
1 | Define tests required to be completed during inspection to ensure modified components are acceptable if needed
Quality Function (During Inspection)
1 | Determine Sample Size of components to be tested based on the Lot Size of shipment
(See Incoming Sampling Process and Escalation)
2 | Test components to newly developed test protocol
3 | Mark the components tested to identify which components received inspection as well as mark boxes to identify
which boxes had the contents inspected
3 | If the components pass inspection follow the Product Release Procedure
4 | If the components fail the inspection follow the Incoming Screening Process and Escalation chart
to determine actions based on testing
5 | Record results of the inspection on 3 time inspect board
Reference Procedures:
1 | Incoming Sampling Process and Escalation (Sample Size Based on MIL-STD-105E)
Determine the Lot size of components
| Lots Size | 1st sample | 2nd sample | 3rd Sample
| 2 to 25 | All | 50% | 5% |  | 1st Inspection Example
| 26 to 50 | 13 | 8 | 5
| 51 to 90 | 20 | 13 | 5 |  | 1000 | Parts
| 91 to 150 | 32 | 20 | 8 |  | 20 | Boxes
| 151 to 280 | 50 | 32 | 13 |  | 125 | Sampled parts
| 281 to 500 | 80 | 50 | 20 |  | 6.25 | Parts per Box
| 501 to 1200 | 125 | 80 | 32
| 1201 to 3200 | 200 | 125 | 50
| 3201 to 10000 | 315 | 200 | 80
· | Complete the sampling based on lot size listed above and increase the sampling effectiveness by selecting
an equal number from each box in the lot based on the sample quantity divided by the number of boxes in the lot size
· | If a defect is found than review the defect with a Quality Engineer to determine if the entire lot can be rejected or
the "Incoming Screening Process and Escalation" below must be followed
2 | Incoming Screening Process and Escalation
1st Inspection | 100% Inspection | Zero Defects | Lot Approved | Moves to 2nd Inspection
|  | Defect < 3% | Review Defects | A Defect | 100% Inspection | Stays at 1st Inspection
|  |  |  | B Defect | Defect caught online | Lot Approved | Moves to 2nd Inspection
[5 empty] | Not caught online | Stays at 1st Inspection
|  |  |  | C Defect | Defect caught online | Lot Approved | Moves to 2nd Inspection
[5 empty] | Not caught online | Stays at 1st Inspection
|  | Defect > 3% | Review Defects | A Defect | 100% Inspection | Stays at 1st Inspection
|  |  |  | B Defect | 100% Inspection | Stays at 1st Inspection
|  |  |  | C Defect | 100% Inspection | Stays at 1st Inspection
2nd Inspection | 50% Inspection | Zero Defects | Lot Approved | Moves to 3rd Inspection
|  | Defect < 3% | Review Defects | A Defect | 100% Inspection | Goes back to 1st Inspection
|  |  |  | B Defect | Defect caught online | Lot Approved | Moves to 3rd Inspection
[5 empty] | Not caught online | Goes back to 1st Inspection
|  |  |  | C Defect | Defect caught online | Lot Approved | Moves to 3rd Inspection
[5 empty] | Not caught online | Stays at 2nd Inspection
|  | Defect > 3% | Review Defects | A Defect | 100% Inspection | Goes back to 1st Inspection
|  |  |  | B Defect | 100% Inspection | Goes back to 1st Inspection
|  |  |  | C Defect | 100% Inspection | Goes back to 1st Inspection
3rd Inspection | 5% Inspection | Zero Defects | Lot Approved | Off Inspection
|  | Defect < 3% | Review Defects | A Defect | 100% Inspection | Goes back to 1st Inspection
|  |  |  | B Defect | Defect caught online | Lot Approved | Stays at 3rd Inspection
[5 empty] | Not caught online | Goes back to 1st Inspection
|  |  |  | C Defect | Defect caught online | Lot Approved | Stays at 3rd Inspection
[5 empty] | Not caught online | Stays at 3rd Inspection
|  | Defect > 3% | Review Defects | A Defect | 100% Inspection | Goes back to 1st Inspection
|  |  |  | B Defect | 100% Inspection | Goes back to 1st Inspection
|  |  |  | C Defect | 100% Inspection | Goes back to 1st Inspection
Duration | 3 Sequential Shipments Defect free or following the above matrix that ends at off inspection
3 | Product Release Procedure
| 1 | Verify part number is located in LC-INSP by Logging into Manage 2000 function LOCS, hit enter once,
and type in INSP / LC-INSP. Once the part number is verified to be in the location end out of LOCS
| 2 | Determine the Purchase Order number that the parts were brought into the facility by using the Manage
2000 function ITH. Type in the part number and find the line number which transferred the parte to
INSP / LC-INSP and record the number for future use and end out of ITH
| 3 | Release the product using Manage 2000 function INSPECT. Enter the Inspection location
(INSP / LC-INSP) enter the Purchase Order Number, Enter the quantity inspected that were good,
Enter the Quantity rejected (if applicable) and enter through the rest of the fields. End out of INSPECT.
| 4 | If product failed inspection contact the commodity / purchasing representative as soon as possible

## Failure Modes 120815
Date: 12/08/2015 (CS)(TD)
NOTE: HHT expects the valve as-received to meet HHT’s technical and performance requirements
ITEM | FAILURE MODES | HHT CONTAINMENT LEVEL
1 | Pilot Fails to Fully Open Within 2 Sec After Receiving Kick Voltage | <3 Flow To Line; >3 Reject Pallet / Lot
2 | Main Fails to Fully Open Within 2 Sec After Receiving Kick Voltage | <3 Flow To Line; >3 Reject Pallet / Lot
3 | Main Fails to Fully Close Within 3 sec of OFF Command | C=0 (Reject Pallet / Lot)
4 | Pilot Fails to Fully Close Within 3 sec of OFF Command | C=0 (Reject Pallet / Lot)
5 | Regulation Knob or Stepper Fails to Regulate HI/LO Adjustment | <3 Flow To Line; >3 Reject Pallet / Lot
6 | Regulation Instability > 0.2 IWC (Oscillation / Pulsating of Pilot or Main) | <3 Flow To Line; >3 Reject Pallet / Lot
7 | Over Pressure (High Flame): Greater Than | C=0 (Reject Pallet / Lot)
750-500: 3.8
750-501: 10.5
2166-302: 3.8
2166-303: 10.5
593-500: 3.8
593-501: 10.5
2155-501: 10.5
8 | Under Pressure (Low Flame): Less Than | C=0 (Reject Pallet / Lot)
750-500: 1.4
750-501: 5.0
2166-302: 0.9
2166-303: 2.6
593-500: 3.2
593-501: 9.5
2155-501: 5.0
9 | Insufficient Output Pressure Regulation Range: Reject If | <3 Flow To Line; >3 Reject Pallet / Lot
750-500: HI < 3.2 or LO >2.0
750-501: HI < 9.5 or LO > 5.5
2166-302: HI < 3.2 or LO > 1.5
2166-303: HI < 9.5 or LO > 3.6
10 | External Leakage | C=0 (Reject Pallet / Lot)
11 | Any combination of non-C=0 defects | <3 Flow To Line; >3 Reject Pallet / Lot

## SW - Valve Insp w Inductance (hidden)
Standard Work - Gas Valve 750-500 Incoming Inspection Process |  |  | Step | Inductance Measurement Procedure
Complete ALL steps for EACH valve, including Fails |  |  | 1 | Use calibrated BK 879B LCR meter. To setup meter for measurements, use BK Inductance Meter Setup tab in Excel file.
Step | Procedure | Results | 2 | With meter set to measure Inductance, touch test leads together, display should read < 2.5 mH.
1 | Randomly select 120 valves for inspection from valve lot, record box label info. |  | 3 | Place one test lead on Main solenoid coil spade and the other on the valve body, when reading is steady, record reading. Acceptable range is 350 to 700 mH, nominal is 480 mH.
1.5 | Pull up the NCM database, and select: NCM Quality Form; "Valves" (under the Sample Inspection section of the Main Form). Once in the inspection screen, choose "Sample Inspection" for Inspection Type. |  | 4 | Place one test lead on Pilot solenoid coil spade and the other on the valve body, when reading is steady, record reading. Acceptable range is 250 to 650 mH, nominal is 380 mH.
2 | Remove valve from box, scan bar code, log valve SN.
3 | Visually inspect valve for missing/broken screws, debris in inlet/outlet ports. Log results. | P/F
4 | Visually inspect valves for indications valve has been dropped, other damage. Log results. | P/F
5 | Measure pilot valve solenoid INDUCTANCE (flag if outside 250-650 mH). Log results. | nom 380 mH
6 | Measure main valve solenoid inductance (flag if outside 350-700 mH). Log results. | nom 480 mH
7 | Install valve on test bench w/pilot & main burner/ECM/RC (if required). Connect manometers to valve pilot & main pressure ports. Set inlet pressure to nominal for valve (NG: 7 IWC, LP: 11 IWC)
8 | With system OFF, use AIM-N-Flame to check for gas leaks at valve and pilot assembly. Log results. | P/F
9 | Turn on pilot/burner with wall switch or RC.
10 | Verify pilot main light within 10 sec. Log results. | P/F
11 | Adjust flame height on RC from high to low to high, confirm flame height adjustment responds properly by manometer reading. Log results. | P/F
12 | Turn off pilot/burner w/wall switch or RC.
13 | With system OFF, use AIM-N-Flame to check for gas leaks at valve and pilot assembly. Log results. | P/F |  | Typical measurement setup
14 | Remove valve from test bench. If pass, place black check mark by valve Date Code, replace in box. If fail, place red sticker by valve Date Code, place valve in Failed Part Tote.
15 | When box is full, mark as QA inspected.
16 | If all valves pass, entire lot is released to production. |  |  | BK 879B LCR Meter
17 | If valve failures, SQE will determine resolution.
If Valve is Dropped during Inspection Process, SCRAP VALVE, DO NOT COUNT AS FAIL, Document in Comment Section in Data Sheet.
Valve in test bench
Manometer Connections
Example Test Bench
Valve solenoids, Stepper Motor, Valve in/out pressure measurement ports

## 2166-302 Stepper Motor SW (hidden)
Test Steps for 2166-302 Stepper Motor Valve
Ensure RC300 is paired to 8K ECM used for valve test. | Put ECM slider switch in REM position, press ECM pairing button, press flame button on RC300, should hear double beep when pairing is successful. | OK?
Ensure Manometer is connected to main test port on valve, ensure inlet pressure is correct for gas used (LP/NG) | 7.5 IWC NG, 10.5 IWC LP standard inlet pressures. | OK?
Turn the FLAME ON with RC300 | The Pilot and Main should ignite (this should happen within 10 sec). Main will light on High (NG 3.5 +/- 0.3, LP 10 +/- 0.5) | Time (Sec)
Adjust Flame level from 5 to 4 with RC300 | Visually see the flame level change and on manometer (NG 2.9 +/- 0.3/LP 8.5+/- 0.5) | IWC?
Adjust Flame level from 4 to 3 with RC300. | Visually see the flame level change and on manometer (NG 2.1 +/- 0.3, LP 6.5 +/- 0.5) | IWC?
Adjust Flame level from 3 to 2 with RC300. | Visually see the flame level change and on manometer (NG 1.4 +/- 0.3, LP 4.5 +/- 0.5) | IWC?
Adjust Flame level from 2 to 1 with RC300. | Visually see the flame level change and on manometer (NG 1.0 +.5/-0.1, LP 2.7 +.09/-.01) | IWC?
Adjust Flame level from 1 to 5 with RC300. | Visually see the flame level change and on manometer (NG 3.5 +/- 0.3, LP 10 +/- 0.5) | IWC?
Switch flame level to 1 with RC300. Shut down the unit and turn back ON with RC. Verify that it lights on HIGH for 10 seconds and toggles down to LOW(1). | Verify unit lights on HIGH and then toggles down to LOW after 10 seconds | YES/NO, Time (sec)
38.6
38
38.7
39
38.5
38.8
38.8
38.8
38.6
38.5
38.3
38.8
38.3
38.4
38.3
38.4
38.5
38.1
38.4
38.2
38.3
38.8
38.5
38.7
39.3
39.2
39
39
39.1
38.9
39
40.9
39.5
39.3
38.7
39.1
38.7
39
39.1
38.9
39.2
39
39
38.7
39
39.4
39
38.6
39.2
39.1
39
39
39.1
39
39.1
39.1
39.2
38.9
39.1
39
39.3
39.3
38.7
39.2
39
39.1
39.2
39
39.1
39.3
39
39.1
38.7
38.8
38.7
38.4
38.6
38.7
38.6
38.8
38.7
38.7
38.9
38.7
39.3
38.5
38.6
38.7
38.6
38.6
39
38.5
38.6
38.8
38.8
38.8
38.7
39.9
38.8
38.7
38.5
38.6
38.6
38.9
38.6
38.5
38.8
38.8
38.7
38.6
38.8
38.6
38.6
38.9
38.7
38.7
38.6
39
38.8
38.9
38.6
38.5
38.8
38.9
38.8
38.6
38.5
38.7
38.8
38.6
38.6
38.8
38.4
38.3
38.6
38.8
38.8
39
38.7
38.8
38.5
38.8
38.6
38.8
38.7
38.2
37.9
38
38.1
38.1
38.6
37.9
38.6
37.9
38.8
38.2
38
38.1
38.3
38.1
38
38.3
38.5
38.8
38
38.5
38.1
38.4
38.5
38.6
38.3
38.5
39
38.3
38.8
38.9
39.1
38.3
39.5
38.4
38.8
38.3
38.4
38.6
38.6
39
39
38.8
38.8
38.8
38.8
38.6
38.1
38.2
37.9
38.5
38.4
38.1
38.5
38.1
38.1
38.2
38.5
37.9
38.2
37.7
38.8
38.3
38.8
38.2
38.7
38.8
38.3
38.5
38.5
38.5
38.5
37.8
38.6
38.9
39.8
38.4
38.8
38.5
39.2
38.7
39.4
39.3
39
38.1
38.3
38.1
38.2
39.7
38
39.5
39.7
38.8
38
39.4
38.2
38.4
38.4
38
37.6
38.1
38.2
37.9
38.3
38.1
38.2
38.6
38.4
39.4
39.4
39
38.2
37.9
39.9
38.3
37.9
39.1
38
39.1
39.3
38.1
38.1
38.8
38.2
38.4
38.5
38.6
38.1
38
38.5
38.2
38.3
38.8
38.6
38.2
38.7
38.2
38.6
38.1
38.8
38.6
38.7
38.5
38.8
38.7
38.8
38.6
39
39.3
38.4
38.4
39.2
38.5
38.6
38.3
38.5
38.4
38.1
38.3
38.2
38.4
38.2
38.1
38.2
38.2
38.3
38.4
38.2
39

## 2166-336 LED Module (hidden)
Test Steps for 2166-336 LED Module
Ensure RC300 is paired to 8K ECM & ACM used for valve test. | Put ECM slider switch in REM position, press ECM pairing button, press flame button on RC300, should hear double beep when pairing is successful. | OK?
Ensure LED module 2166-336 is connected to AUX1 port of ACM and LED light tubs | check connections | OK?
Turn AUX1 output on HIGH (all three bars) with RC300 | Visually see the LED light level change to HIGH | OK?
Turn AUX1 output on MED (two bars) with RC300 | Visually see the LED light level change to MEDIUM | OK?
Turn AUX1 output on OFF with RC300 | Visually see the LED light turn OFF | OK?

## Valve Data Sheet (hidden)
Valve Bar Code/SN | Step 3 Results - Visual Insp, damaged screws/debris | Step 4 Results - Drop/impact damge check | Step 5 Results - Pilot Valve Solenoid Resistance (ohms) | Step 6 Results - Main Valve Solenoid Resistance (ohms) | Step 8 Results - Leak Test of valve and pilot | Step 10 Results - Main lights in <10 sec | Step 11 Results - Flame Height Adjustment Works | Step 13 Results - Leak Test of valve and pilot | Comments
Note on A1: Jeff Sulla: Bar code scanned for part no/serial no of each valve. Use cheap $90 USB bar code scanner at each station to fill EXCEL spreadsheet. Inspector will complete cells in each row for each valve inspected (P - pass, F - fail).

## Valve Pereto 1-15-15 (hidden)
# | Date Analyzed | Date Reported | Part # | Main Inductance (Henry) | Main Resistance (Ohms) | Dexen Insp # | Grand Mate Insp # | Gas Type | Problem | SCAR | RMA # | Shipped | Return to Dexen | Analyze @ HHT | Return to Grand Mate | Notes | Customer RMA #
1 | 42020 | 41978 | NGK-DXV-50 | 0.662 | 59.4 | 152 | 0961 | NG | N/A | YES |  |  | YES |  |  | Dissasembled | 2604603
2 | 42023 | 41996 | NGK-DXV-50 | 0.687 | 59.1 | 152 | 0961 | NG | Does not regulate up or down | YES |  |  | YES
3 | 42023 | 41967 | NGK-DXV-50 | 0.595 | 55 | 158 | NA | NG |  | YES |  |  | YES |  |  |  | 2595525
4 | 42023 | 41961 | NGK-DXV-50 | 0.593 | 53 | NA | NA | NG | Flame height does not change, can hear stepper motor move | YES |  |  | YES |  |  | Black wire pulled out of stepper motor | 2589277
5 | 42023 | 41969 | NGK-DXV-50 | 0.513 | 56.4 | NA | NA | NG |  | YES |  |  | YES |  |  |  | 2597506
6 | 42023 | 41999 | NGK-DXV-50 | OL | OL | 152 | 0961 | NG |  | YES |  |  | YES |  |  | Dissassembled, Broken coil wire | 2626309
7 | 42023 | NA | LPK-DXF | 0.566 | 59.7 | 117 | 0961 | LP |  | YES |  |  | YES
8 | 42023 | NA | NGK-DXV-50 | 0.585 | 59.2 | 149 | NA | NG | Stepper motor stuck on low | YES |  |  | YES |  |  |  | 2606023
9 | 42023 | 41978 | NGK-DXV-50 | 0.501 | 58.7 | 158 | NA | NG |  | YES |  |  | YES |  |  |  | 2605305
10 | 42023 | 41985 | NGK-DXV-50 | 0.532 | 59.7 | NA | 0961 | NG | Noisey stepper motor | YES |  |  | YES |  |  |  | 2612154
11 | 42023 | 41990 | LPK-DXV-50 | 0.77 | 56.7 | 152 | 0961 | LP | Defective | YES |  |  | YES |  |  |  | 260275
12 | 42023 | NA | NGK-DXV-50 | 0.635 | 55.9 | 158 | NA | NG |  | YES |  |  | YES |  |  | 1 screw missing | 2590900
13 | 42023 | 41979 | LPK-DXV-50 | 0.271 | 53.5 | 152 | 0961 | LP |  | YES |  |  | YES |  |  |  | 2606014
14 | 42023 | 41979 | LPK-DXV-50 | 0.564 | 50 | 117 | 0961 | LP |  | YES |  |  | YES |  |  | SS screws, resistance is jumpy | 2606014

## C=0 Sample Size Selector (hidden)
Enter Blue cell data | C=0 Sample plan
% Reliability (lower limit) | Confidence Level % | N (sample size)
95 | 99.9945 | 192
.
192

## Valve Failure Modes (hidden)
Valve Failure Modes | Detection
Pilot Solenoid Ohms out of spec | resistance measurement
Main Solenoid Ohms out of spec | resistance measurement
Pilot Valve Stuck Shut | With ON command, no gas flow to pilot burner
Pilot Valve Stuck Open | With OFF command, continuous gas flow to pilot burner
Main Valve Stuck Shut | With ON command, no gas flow to main burner
Main Valve Stuck Open | With OFF command, continuous gas flow to main burner
No Flame Height Adjustment | Visual flame observation or no change in manometer reading with flame height adjustment
Gas leak | Sustained flame with aim & flame
Broken/missing screws | Visual observation/inspection
Broken electrical connectors on solenoids | Visual observation/inspection